## EU-RoHS-AnnexIV_20220131
Exemption List Identification |  |  | IEC 62474 Identification |  |  |  | Information from Regulation |  |  |  |  | Additional Information | [6 empty] | Regulatory Basis
List Authority | List Identity | List Version | IEC-ID format: nnnnn-v-ss, where nnnnn=ID number v=version number ss-subpart | base ID number | Entry version (A, B, C, …) | Subidentity (for portion of exemption (sub-exemption) split up based on expiry or technical requirements) subidentity="00" refers to the full exemption as per regulation | regIndex (Exemption number in Regulation if specified) | Description (Exemption Description) | Scope and dates of applicability | Details2 (if applicable) | Details3 (if applicable) | Status (Informative) (based on date of posting) A=Active E=Expired F=Future | Product Categories | Application (e.g. product parts) | Substance name | DSL ID | Effective Date | Expiry Date | Basis of Substance Regulation | Basis of Exemption Regulation | First Added | Last Revised | Comments
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01001-A-00 | 01001 | A | 00 | 1 | Equipment utilising or detecting ionising radiation; Lead, cadmium and mercury in detectors for ionising radiation. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01001-A-01 | 01001 | A | 01 | 1 | Equipment utilising or detecting ionising radiation; Lead, cadmium and mercury in detectors for ionising radiation. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-02 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01001-A-02 | 01001 | A | 02 | 1 | Equipment utilising or detecting ionising radiation; Lead, cadmium and mercury in detectors for ionising radiation. |  |  |  | Active | 8 in vitro |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01001-A-03 | 01001 | A | 03 | 1 | Equipment utilising or detecting ionising radiation; Lead, cadmium and mercury in detectors for ionising radiation. |  |  |  | Active | 9 industrial |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-06 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01002-A-00 | 01002 | A | 00 | 2 | Equipment utilising or detecting ionising radiation; Lead bearings in X-ray tubes. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01002-A-01 | 01002 | A | 01 | 2 | Equipment utilising or detecting ionising radiation; Lead bearings in X-ray tubes. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-15 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01002-A-02 | 01002 | A | 02 | 2 | Equipment utilising or detecting ionising radiation; Lead bearings in X-ray tubes. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01002-A-03 | 01002 | A | 03 | 2 | Equipment utilising or detecting ionising radiation; Lead bearings in X-ray tubes. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01003-A-00 | 01003 | A | 00 | 3 | Equipment utilising or detecting ionising radiation; Lead in electromagnetic radiation amplification devices: micro-channel plate and capillary plate. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01003-A-01 | 01003 | A | 01 | 3 | Equipment utilising or detecting ionising radiation; Lead in electromagnetic radiation amplification devices: micro-channel plate and capillary plate. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01003-A-02 | 01003 | A | 02 | 3 | Equipment utilising or detecting ionising radiation; Lead in electromagnetic radiation amplification devices: micro-channel plate and capillary plate. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01003-A-03 | 01003 | A | 03 | 3 | Equipment utilising or detecting ionising radiation; Lead in electromagnetic radiation amplification devices: micro-channel plate and capillary plate. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01004-A-00 | 01004 | A | 00 | 4 | Equipment utilising or detecting ionising radiation; Lead in glass frit of X-ray tubes and image intensifiers and lead in glass frit binder for assembly of gas lasers and for vacuum tubes that convert electromagnetic radiation into electrons. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01004-A-01 | 01004 | A | 01 | 4 | Equipment utilising or detecting ionising radiation; Lead in glass frit of X-ray tubes and image intensifiers and lead in glass frit binder for assembly of gas lasers and for vacuum tubes that convert electromagnetic radiation into electrons. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01004-A-02 | 01004 | A | 02 | 4 | Equipment utilising or detecting ionising radiation; Lead in glass frit of X-ray tubes and image intensifiers and lead in glass frit binder for assembly of gas lasers and for vacuum tubes that convert electromagnetic radiation into electrons. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01004-A-03 | 01004 | A | 03 | 4 | Equipment utilising or detecting ionising radiation; Lead in glass frit of X-ray tubes and image intensifiers and lead in glass frit binder for assembly of gas lasers and for vacuum tubes that convert electromagnetic radiation into electrons. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01005-A-00 | 01005 | A | 00 | 5 | Equipment utilising or detecting ionising radiation; Lead in shielding for ionising radiation. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01005-A-01 | 01005 | A | 01 | 5 | Equipment utilising or detecting ionising radiation; Lead in shielding for ionising radiation. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-06 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01005-A-02 | 01005 | A | 02 | 5 | Equipment utilising or detecting ionising radiation; Lead in shielding for ionising radiation. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01005-A-03 | 01005 | A | 03 | 5 | Equipment utilising or detecting ionising radiation; Lead in shielding for ionising radiation. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-06 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01006-A-00 | 01006 | A | 00 | 6 | Equipment utilising or detecting ionising radiation; Lead in X-ray test objects. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01006-A-01 | 01006 | A | 01 | 6 | Equipment utilising or detecting ionising radiation; Lead in X-ray test objects. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01006-A-02 | 01006 | A | 02 | 6 | Equipment utilising or detecting ionising radiation; Lead in X-ray test objects. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01006-A-03 | 01006 | A | 03 | 6 | Equipment utilising or detecting ionising radiation; Lead in X-ray test objects. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01007-A-00 | 01007 | A | 00 | 7 | Equipment utilising or detecting ionising radiation; Lead stearate X-ray diffraction crystals. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01007-A-01 | 01007 | A | 01 | 7 | Equipment utilising or detecting ionising radiation; Lead stearate X-ray diffraction crystals. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01007-A-02 | 01007 | A | 02 | 7 | Equipment utilising or detecting ionising radiation; Lead stearate X-ray diffraction crystals. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01007-A-03 | 01007 | A | 03 | 7 | Equipment utilising or detecting ionising radiation; Lead stearate X-ray diffraction crystals. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01008-A-00 | 01008 | A | 00 | 8 | Equipment utilising or detecting ionising radiation; Radioactive cadmium isotope source for portable X-ray fluorescence spectrometers. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Cadmium | 00010 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01008-A-01 | 01008 | A | 01 | 8 | Equipment utilising or detecting ionising radiation; Radioactive cadmium isotope source for portable X-ray fluorescence spectrometers. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Cadmium | 00010 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01008-A-02 | 01008 | A | 02 | 8 | Equipment utilising or detecting ionising radiation; Radioactive cadmium isotope source for portable X-ray fluorescence spectrometers. |  |  |  | Active | 8 in vitro |  | Cadmium | 00010 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01008-A-03 | 01008 | A | 03 | 8 | Equipment utilising or detecting ionising radiation; Radioactive cadmium isotope source for portable X-ray fluorescence spectrometers. |  |  |  | Active | 9 industrial |  | Cadmium | 00010 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01009-A-00 | 01009 | A | 00 | 1a | Sensors, detectors and electrodes; Lead and cadmium in ion selective electrodes including glass of pH electrodes. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01009-A-01 | 01009 | A | 01 | 1a | Sensors, detectors and electrodes; Lead and cadmium in ion selective electrodes including glass of pH electrodes. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-17 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01009-A-02 | 01009 | A | 02 | 1a | Sensors, detectors and electrodes; Lead and cadmium in ion selective electrodes including glass of pH electrodes. |  |  |  | Active | 8 in vitro |  | Cadmium; Lead | 00010; 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01009-A-03 | 01009 | A | 03 | 1a | Sensors, detectors and electrodes; Lead and cadmium in ion selective electrodes including glass of pH electrodes. |  |  |  | Active | 9 industrial |  | Cadmium; Lead | 00010; 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-17 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01010-A-00 | 01010 | A | 00 | 1b | Sensors, detectors and electrodes; Lead anodes in electrochemical oxygen sensors. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01010-A-01 | 01010 | A | 01 | 1b | Sensors, detectors and electrodes; Lead anodes in electrochemical oxygen sensors. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01010-A-02 | 01010 | A | 02 | 1b | Sensors, detectors and electrodes; Lead anodes in electrochemical oxygen sensors. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01010-A-03 | 01010 | A | 03 | 1b | Sensors, detectors and electrodes; Lead anodes in electrochemical oxygen sensors. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01011-A-00 | 01011 | A | 00 | 1c | Sensors, detectors and electrodes; Lead, cadmium and mercury in infra-red light detectors. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01011-A-01 | 01011 | A | 01 | 1c | Sensors, detectors and electrodes; Lead, cadmium and mercury in infra-red light detectors. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2019-11-22 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01011-A-02 | 01011 | A | 02 | 1c | Sensors, detectors and electrodes; Lead, cadmium and mercury in infra-red light detectors. |  |  |  | Active | 8 in vitro |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01011-A-03 | 01011 | A | 03 | 1c | Sensors, detectors and electrodes; Lead, cadmium and mercury in infra-red light detectors. |  |  |  | Active | 9 industrial |  | Cadmium; Lead; Mercury | 00010; 00021; 00029 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2019-12-23 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01012-A-00 | 01012 | A | 00 | 1d | Sensors, detectors and electrodes; Mercury in reference electrodes: low chloride mercury chloride, mercury sulphate and mercury oxide. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Mercury | 00029 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01012-A-01 | 01012 | A | 01 | 1d | Sensors, detectors and electrodes; Mercury in reference electrodes: low chloride mercury chloride, mercury sulphate and mercury oxide. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Mercury | 00029 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01012-A-02 | 01012 | A | 02 | 1d | Sensors, detectors and electrodes; Mercury in reference electrodes: low chloride mercury chloride, mercury sulphate and mercury oxide. |  |  |  | Active | 8 in vitro |  | Mercury | 00029 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01012-A-03 | 01012 | A | 03 | 1d | Sensors, detectors and electrodes; Mercury in reference electrodes: low chloride mercury chloride, mercury sulphate and mercury oxide. |  |  |  | Active | 9 industrial |  | Mercury | 00029 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01013-A-00 | 01013 | A | 00 | 9 | Others; Cadmium in helium-cadmium lasers. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Cadmium | 00010 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01013-A-01 | 01013 | A | 01 | 9 | Others; Cadmium in helium-cadmium lasers. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Cadmium | 00010 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01013-A-02 | 01013 | A | 02 | 9 | Others; Cadmium in helium-cadmium lasers. |  |  |  | Active | 8 in vitro |  | Cadmium | 00010 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01013-A-03 | 01013 | A | 03 | 9 | Others; Cadmium in helium-cadmium lasers. |  |  |  | Active | 9 industrial |  | Cadmium | 00010 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01014-A-00 | 01014 | A | 00 | 10 | Others; Lead and cadmium in atomic absorption spectroscopy lamps. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01014-A-01 | 01014 | A | 01 | 10 | Others; Lead and cadmium in atomic absorption spectroscopy lamps. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01014-A-02 | 01014 | A | 02 | 10 | Others; Lead and cadmium in atomic absorption spectroscopy lamps. |  |  |  | Active | 8 in vitro |  | Cadmium; Lead | 00010; 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01014-A-03 | 01014 | A | 03 | 10 | Others; Lead and cadmium in atomic absorption spectroscopy lamps. |  |  |  | Active | 9 industrial |  | Cadmium; Lead | 00010; 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01015-A-00 | 01015 | A | 00 | 11 | Others; Lead in alloys as a superconductor and thermal conductor in MRI. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01015-A-01 | 01015 | A | 01 | 11 | Others; Lead in alloys as a superconductor and thermal conductor in MRI. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-02 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01015-A-02 | 01015 | A | 02 | 11 | Others; Lead in alloys as a superconductor and thermal conductor in MRI. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01015-A-03 | 01015 | A | 03 | 11 | Others; Lead in alloys as a superconductor and thermal conductor in MRI. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01016-A-00 | 01016 | A | 00 | 12 | Others; Lead and cadmium in metallic bonds to superconducting materials in MRI and SQUID detectors. | Expired on: 28 January 2014 |  |  | Expired | 8 and 9 |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | 2014-01-28 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | In Annex IV to Directive 2011/65/EU point 12 is replaced by the following: See Description (Exemption Description) of IEC-ID 01016-B-00. E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01016-B-00 | 01016 | B | 00 | 12 | Others; Lead and cadmium in metallic bonds creating superconducting magnetic circuits in MRI, SQUID, NMR (Nuclear Magnetic Resonance) or FTMS (Fourier Transform Mass Spectrometer) detectors. | Expires on: 30 June 2021. |  |  | Active | 8 and 9 |  | Cadmium; Lead | 00010; 00021 | 2014-7-22 | TBD | 2011/65/EU | 2014/9/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2019-07-31(Partially) E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01017-A-00 | 01017 | A | 00 | 13 | Others; Lead in counterweights. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01017-A-01 | 01017 | A | 01 | 13 | Others; Lead in counterweights. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01017-A-02 | 01017 | A | 02 | 13 | Others; Lead in counterweights. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01017-A-03 | 01017 | A | 03 | 13 | Others; Lead in counterweights. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01018-A-00 | 01018 | A | 00 | 14 | Others; Lead in single crystal piezoelectric materials for ultrasonic transducers. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01018-A-01 | 01018 | A | 01 | 14 | Others; Lead in single crystal piezoelectric materials for ultrasonic transducers. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-02 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01018-A-02 | 01018 | A | 02 | 14 | Others; Lead in single crystal piezoelectric materials for ultrasonic transducers. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01018-A-03 | 01018 | A | 03 | 14 | Others; Lead in single crystal piezoelectric materials for ultrasonic transducers. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01019-A-00 | 01019 | A | 00 | 15 | Others; Lead in solders for bonding to ultrasonic transducers. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01019-A-01 | 01019 | A | 01 | 15 | Others; Lead in solders for bonding to ultrasonic transducers. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01019-A-02 | 01019 | A | 02 | 15 | Others; Lead in solders for bonding to ultrasonic transducers. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01019-A-03 | 01019 | A | 03 | 15 | Others; Lead in solders for bonding to ultrasonic transducers. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01020-A-00 | 01020 | A | 00 | 16 | Others; Mercury in very high accuracy capacitance and loss measurement bridges and in high frequency RF switches and relays in monitoring and control instruments not exceeding 20 mg of mercury per switch or relay. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Mercury | 00029 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01020-A-01 | 01020 | A | 01 | 16 | Others; Mercury in very high accuracy capacitance and loss measurement bridges and in high frequency RF switches and relays in monitoring and control instruments not exceeding 20 mg of mercury per switch or relay. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Mercury | 00029 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01020-A-02 | 01020 | A | 02 | 16 | Others; Mercury in very high accuracy capacitance and loss measurement bridges and in high frequency RF switches and relays in monitoring and control instruments not exceeding 20 mg of mercury per switch or relay. |  |  |  | Active | 8 in vitro |  | Mercury | 00029 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01020-A-03 | 01020 | A | 03 | 16 | Others; Mercury in very high accuracy capacitance and loss measurement bridges and in high frequency RF switches and relays in monitoring and control instruments not exceeding 20 mg of mercury per switch or relay. |  |  |  | Active | 9 industrial |  | Mercury | 00029 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01021-A-00 | 01021 | A | 00 | 17 | Others; Lead in solders in portable emergency defibrillators. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01021-A-01 | 01021 | A | 01 | 17 | Others; Lead in solders in portable emergency defibrillators. |  |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-15 1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01021-A-02 | 01021 | A | 02 | 17 | Others; Lead in solders in portable emergency defibrillators. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01021-A-03 | 01021 | A | 03 | 17 | Others; Lead in solders in portable emergency defibrillators. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01022-A-00 | 01022 | A | 00 | 18 | Others; Lead in solders of high performance infrared imaging modules to detect in the range 8-14 μm. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01022-A-01 | 01022 | A | 01 | 18 | Others; Lead in solders of high performance infrared imaging modules to detect in the range 8-14 μm. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01022-A-02 | 01022 | A | 02 | 18 | Others; Lead in solders of high performance infrared imaging modules to detect in the range 8-14 μm. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01022-A-03 | 01022 | A | 03 | 18 | Others; Lead in solders of high performance infrared imaging modules to detect in the range 8-14 μm. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01023-A-00 | 01023 | A | 00 | 19 | Others; Lead in Liquid crystal on silicon (LCoS) displays. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01023-A-01 | 01023 | A | 01 | 19 | Others; Lead in Liquid crystal on silicon (LCoS) displays. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01023-A-02 | 01023 | A | 02 | 19 | Others; Lead in Liquid crystal on silicon (LCoS) displays. |  |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01023-A-03 | 01023 | A | 03 | 19 | Others; Lead in Liquid crystal on silicon (LCoS) displays. |  |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01024-A-00 | 01024 | A | 00 | 20 | Others; Cadmium in X-ray measurement filters. | Expires on: 7-years exemption applies to medical devices and monitoring and control instruments which are placed on the market from 22 July 2014, to in vitro diagnostic medical devices which are placed on the market from 22 July 2016 and to industrial monitoring and control instruments which are placed on the market from 22 July 2017. |  |  | Partially expired | 8 and 9 |  | Cadmium | 00010 | 2014-7-22 | Multiple | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption partially expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01024-A-01 | 01024 | A | 01 | 20 | Others; Cadmium in X-ray measurement filters. |  |  |  | Expired | 8 and 9 other than in vitro and industrial |  | Cadmium | 00010 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01024-A-02 | 01024 | A | 02 | 20 | Others; Cadmium in X-ray measurement filters. |  |  |  | Active | 8 in vitro |  | Cadmium | 00010 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01024-A-03 | 01024 | A | 03 | 20 | Others; Cadmium in X-ray measurement filters. |  |  |  | Active | 9 industrial |  | Cadmium | 00010 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2011/65/EU | 2019-11-30 | 2019-11-30 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01025-A-00 | 01025 | A | 00 | 21 | Others; Cadmium in phosphor coatings in image intensifiers for X-ray images until 31 December 2019 and in spare parts for X-ray systems placed on the EU market before 1 January 2020. | Expires on: 31 December 2019 and in spare parts for X-ray systems placed on the EU market before 1 January 2020. |  |  | Expired | 8 and 9 |  | Cadmium | 00010 | 2014-7-22 | 2019-12-31 | 2011/65/EU | 2014/2/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2019-12-31; Editorial Corrections to Expiry Date E2.0: Exemption Status update 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01026-A-00 | 01026 | A | 00 | 22 | Others; Lead acetate marker for use in stereotactic head frames for use with CT and MRI and in positioning systems for gamma beam and particle therapy equipment. | Expires on: 30 June 2021 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2021-06-30 | 2011/65/EU | 2014/3/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-06-30 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01027-A-00 | 01027 | A | 00 | 23 | Others; Lead as an alloying element for bearings and wear surfaces in medical equipment exposed to ionising radiation. | Expires on: 30 June 2021 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2021-06-30 | 2011/65/EU | 2014/1/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-06-30 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01028-A-00 | 01028 | A | 00 | 24 | Others; Lead enabling vacuum tight connections between aluminium and steel in X-ray image intensifiers. | Expires on: 31 December 2019 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2019-12-31 | 2011/65/EU | 2014/4/EU | 2019-11-30 | 2021-01-19 | E2.0: Exemption expired 2019-12-31 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01029-A-00 | 01029 | A | 00 | 25 | Others; Lead in the surface coatings of pin connector systems requiring nonmagnetic connectors which are used durably at a temperature below – 20 °C under normal operating and storage conditions | Expires on: 30 June 2021 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2021-06-30 | 2011/65/EU | 2014/6/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-06-30 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01030-A-00 | 01030 | A | 00 | 26 | Others; Lead in — solders on printed circuit boards, — termination coatings of electrical and electronic components and coatings of printed circuit boards, — solders for connecting wires and cables, — solders connecting transducers and sensors, that are used durably at a temperature below – 20 °C under normal operating and storage conditions. | Expired on: 14 July 2016 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2016-07-14 | 2011/65/EU | 2014/5/EU | 2019-11-30 | 2021-01-19 | In Annex IV to Directive 2011/65/EU, point 26 is replaced by the following: See Description (Exemption Description) of IEC-ID 01030-B-00. E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01030-B-00 | 01030 | B | 00 | 26 | Others; Lead in the following applications that are used durably at a temperature below – 20 °C under normal operating and storage conditions: (a) solders on printed circuit boards; (b) termination coatings of electrical and electronic components and coatings of printed circuit boards; (c) solders for connecting wires and cables; (d) solders connecting transducers and sensors. Lead in solders of electrical connections to temperature measurement sensors in devices which are designed to be used periodically at temperatures below – 150 °C. | Expires on: 30 June 2021 |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2016/1028/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2019-12-11 E2.0: Editorial Corrections to Basis of Exemption Regulation 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01031-A-00 | 01031 | A | 00 | 27 | Others; Lead in — solders, — termination coatings of electrical and electronic components and printed circuit boards, — connections of electrical wires, shields and enclosed connectors, which are used in (a) magnetic fields within the sphere of 1 m radius around the isocentre of the magnet in medical magnetic resonance imaging equipment, including patient monitors designed to be used within this sphere, or (b) magnetic fields within 1 m distance from the external surfaces of cyclotron magnets, magnets for beam transport and beam direction control applied for particle therapy. | Expires on: 30 June 2020 |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2014/7/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 20１8-12-12(Partially) E2.0: Exemption Status update 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01032-A-00 | 01032 | A | 00 | 28 | Others; Lead in solders for mounting cadmium telluride and cadmium zinc telluride digital array detectors to printed circuit boards. | Expired on: 31 December 2017 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2017-12-31 | 2011/65/EU | 2014/8/EU | 2019-11-30 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01033-A-00 | 01033 | A | 00 | 29 | Others; Lead in alloys, as a superconductor or thermal conductor, used in cryo-cooler cold heads and/or in cryo-cooled cold probes and/or in cryo-cooled equipotential bonding systems, in medical devices (category 8) and/or in industrial monitoring and control instruments. | Expires on: 30 June 2021 |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2014/10/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2019-12-20(Partially) E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01034-A-00 | 01034 | A | 00 | 30 | Others; Hexavalent chromium in alkali dispensers used to create photocathodes in X-ray image intensifiers until 31 December 2019 and in spare parts for X-ray systems placed on the EU market before 1 January 2020. | Expires on: 31 December 2019 and in spare parts for X-ray systems placed on the EU market before 1 January 2020. |  |  | Expired | 8 and 9 |  | Hexavalent Chromium | 00012 | 2014-7-22 | 2019-12-31 | 2011/65/EU | 2014/11/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2019-12-31 E2.0: Exemption Status update 2021-01-19
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01035-A-00 | 01035 | A | 00 | 31 | Others; Lead, cadmium and hexavalent chromium in reused spare parts, recovered from medical devices placed on the market before 22 July 2014 and used in category 8 equipment placed on the market before 22 July 2021, provided that reuse takes place in auditable closed-loop business-to-business return systems, and that the reuse of parts is notified to the consumer. | Expired on: 5 November 2017 |  |  | Expired | 8 and 9 |  | Cadmium; Hexavalent Chromium; Lead | 00010; 00012; 00021 | 2014-7-22 | 2017-11-05 | 2011/65/EU | 2014/15/EU | 2019-11-30 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01035-B-00 | 01035 | B | 00 | 31a | Others; Lead, cadmium, hexavalent chromium, and polybrominated diphenyl ethers (PBDE) in spare parts recovered from and used for the repair or refurbishment of medical devices, including in vitro diagnostic medical devices, or electron microscopes and their accessories, provided that the reuse takes place in auditable closed-loop business-to-business return systems and that each reuse of parts is notified to the customer. | Expires on:(a) 21 July 2021 for the use in medical devices other than in vitro diagnostic medical devices;(b) 21 July 2023 for the use in in vitro diagnostic medical devices;(c) 21 July 2024 for the use in electron microscopes and their accessories. |  |  | Active | 8 and 9 |  | Cadmium; Hexavalent Chromium; Lead; PBDE | 00010; 00012; 00021; 00045 | 2017-11-06 | Multiple | 2011/65/EU | 2016/585/EU | 2019-11-30 | 2021-01-19 | E2.0: Editorial Corrections to Basis of Exemption Regulation 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01035-B-01 | 01035 | B | 01 | 31a | Others; Lead, cadmium, hexavalent chromium, and polybrominated diphenyl ethers (PBDE) in spare parts recovered from and used for the repair or refurbishment of medical devices, including in vitro diagnostic medical devices, or electron microscopes and their accessories, provided that the reuse takes place in auditable closed-loop business-to-business return systems and that each reuse of parts is notified to the customer. | Expires on:(a) 21 July 2021 for the use in medical devices other than in vitro diagnostic medical devices; |  |  | Active | 8 and 9 other than in vitro and industrial |  | Cadmium; Hexavalent Chromium; Lead; PBDE | 00010; 00012; 00021; 00045 | 2017-11-06 | TBD | 2011/65/EU | 2016/585/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Editorial Corrections to Product Categories; Request for renewal 2020-01-02 E2.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01035-B-02 | 01035 | B | 02 | 31a | Others; Lead, cadmium, hexavalent chromium, and polybrominated diphenyl ethers (PBDE) in spare parts recovered from and used for the repair or refurbishment of medical devices, including in vitro diagnostic medical devices, or electron microscopes and their accessories, provided that the reuse takes place in auditable closed-loop business-to-business return systems and that each reuse of parts is notified to the customer. | Expires on:(b) 21 July 2023 for the use in in vitro diagnostic medical devices; |  |  | Active | 8 in vitro |  | Cadmium; Hexavalent Chromium; Lead; PBDE | 00010; 00012; 00021; 00045 | 2017-11-06 | 2023-07-21 | 2011/65/EU | 2016/585/EU | 2019-11-30 | 2021-01-19 | E2.0: Editorial Corrections to Basis of Exemption Regulation 2021-01-19 E1.0: Added exemption; Request for renewal 2020-01-02
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01035-B-03 | 01035 | B | 03 | 31a | Others; Lead, cadmium, hexavalent chromium, and polybrominated diphenyl ethers (PBDE) in spare parts recovered from and used for the repair or refurbishment of medical devices, including in vitro diagnostic medical devices, or electron microscopes and their accessories, provided that the reuse takes place in auditable closed-loop business-to-business return systems and that each reuse of parts is notified to the customer. | Expires on:(c) 21 July 2024 for the use in electron microscopes and their accessories. |  |  | Active | 9 |  | Cadmium; Hexavalent Chromium; Lead; PBDE | 00010; 00012; 00021; 00045 | 2017-11-06 | 2024-07-21 | 2011/65/EU | 2016/585/EU | 2019-11-30 | 2021-01-19 | E2.0: Editorial Corrections to Basis of Exemption Regulation 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01036-A-00 | 01036 | A | 00 | 32 | Others; Lead in solders on printed circuit boards of detectors and data acquisition units for Positron Emission Tomographs which are integrated into Magnetic Resonance Imaging equipment. | Expires on: 31 December 2019 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2019-12-31 | 2011/65/EU | 2014/12/EU | 2019-11-30 | 2021-01-19 | E2.0: Exemption expired 2019-12-31 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01037-A-00 | 01037 | A | 00 | 33 | Others; Lead in solders on populated printed circuit boards used in Directive 93/42/EEC class IIa and IIb mobile medical devices other than portable emergency defibrillators. | Expires on: 30 June 2016 for class IIa and on 31 December 2020 for class IIb. |  |  | Partially Expired | 8 and 9 |  | Lead | 00021 | 2014-07-22 | Multiple | 2011/65/EU | 2014/13/EU | 2019-11-30 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01037-A-01 | 01037 | A | 01 | 33 | Others; Lead in solders on populated printed circuit boards used in Directive 93/42/EEC class IIa and IIb mobile medical devices other than portable emergency defibrillators. | Expires on: 30 June 2016 for class IIa. |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-07-22 | 2016-06-30 | 2011/65/EU | 2014/13/EU | 2019-11-30 | 2022-01-31 | E4.0: Editorial correction to Scope and dates of applicability E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01037-A-02 | 01037 | A | 02 | 33 | Others; Lead in solders on populated printed circuit boards used in Directive 93/42/EEC class IIa and IIb mobile medical devices other than portable emergency defibrillators. | Expires on: 31 December 2020 for class IIb. |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-07-22 | 2020-12-31 | 2011/65/EU | 2014/13/EU | 2019-11-30 | 2022-01-31 | E4.0: Editorial correction to Scope and dates of applicability E2.0: Exemption expired 2020-12-31 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01038-A-00 | 01038 | A | 00 | 34 | Others; Lead as an activator in the fluorescent powder of discharge lamps when used for extracorporeal photopheresis lamps containing BSP (BaSi2O5:Pb) phosphors. | Expires on: 22 July 2021 |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2021-07-21 | 2011/65/EU | 2014/16/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2021-07-21 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01039-A-00 | 01039 | A | 00 | 35 | Others; Mercury in cold cathode fluorescent lamps for back-lighting liquid crystal displays, not exceeding 5 mg per lamp, used in industrial monitoring and control instruments placed on the market before 22 July 2017. | Expires on: 21 July 2024 |  |  | Active | 9 industrial |  | Mercury | 00029 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2014/75/EU | 2019-11-30 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01040-A-00 | 01040 | A | 00 | 36 | Others; Lead used in other than C-press compliant pin connector systems for industrial monitoring and control instruments. | Expires on: 31 December 2020. May be used after that date in spare parts for industrial monitoring and control instruments placed on the market before 1 January 2021. |  |  | Expired | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2020-12-31 | 2011/65/EU | 2014/74/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2020-12-31 E2.0: Exemption Status update 2021-01-19
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01041-A-00 | 01041 | A | 00 | 37 | Others; Lead in platinized platinum electrodes used for conductivity measurements where at least one of the following conditions applies: (a) wide-range measurements with a conductivity range covering more than 1 order of magnitude (e.g. range between 0,1 mS/m and 5 mS/m) in laboratory applications for unknown concentrations; (b) measurements of solutions where an accuracy of +/– 1 % of the sample range and where high corrosion resistance of the electrode are required for any of the following: (i) solutions with an acidity < pH 1; (ii) solutions with an alkalinity > pH 13; (iii) corrosive solutions containing halogen gas; (c) measurements of conductivities above 100 mS/m that must be performed with portable instruments. | Expires on: 31 December 2018. |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2018-12-31 | 2011/65/EU | 2014/73/EU | 2019-11-30 | 2021-01-19 | E2.0: Exemption replaced and expired 2018-12-31 In Annex IV to Directive 2011/65/EU, entry 37 is replaced by the following: See Description (Exemption Description) of IEC-ID 01041-B-00. E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01041-B-00 | 01041 | B | 00 | 37 | Others; Lead in platinized platinum electrodes used for conductivity measurements where at least one of the following conditions applies: (a) wide-range measurements with a conductivity range covering more than 1 order of magnitude (e.g. range between 0,1 mS/m and 5 mS/m) in laboratory applications for unknown concentrations; (b) measurements of solutions where an accuracy of +/– 1 % of the sample range and where high corrosion resistance of the electrode are required for any of the following: (i) solutions with an acidity < pH 1; (ii) solutions with an alkalinity > pH 13; (iii) corrosive solutions containing halogen gas; (c) measurements of conductivities above 100 mS/m that must be performed with portable instruments. | Expires on 31 December 2025 |  |  | Active | 8 and 9 |  | Lead | 00021 | 2018-12-31 | 2025-12-31 | 2011/65/EU | 2020/360/EU | 2021-01-19 | 2021-01-19 | E2.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01042-A-00 | 01042 | A | 00 | 38 | Others; Lead in solder in one interface of large area stacked die elements with more than 500 interconnects per interface which are used in X-ray detectors of computed tomography and X-ray systems. | Expires on: 31 December 2019. May be used after that date in spare parts for CT and X-ray systems placed on the market before 1 January 2020. |  |  | Expired | 8 and 9 |  | Lead | 00021 | 2014-7-22 | 2019-12-31 | 2011/65/EU | 2014/71/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2019-12-31 E2.0: Exemption Status update 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01043-A-00 | 01043 | A | 00 | 39 | Others; Lead in micro-channel plates (MCPs) used in equipment where at least one of the following properties is present: (a) a compact size of the detector for electrons or ions, where the space for the detector is limited to a maximum of 3 mm/MCP (detector thickness + space for installation of the MCP), a maximum of 6 mm in total, and an alternative design yielding more space for the detector is scientifically and technically impracticable; (b) a two-dimensional spatial resolution for detecting electrons or ions, where at least one of the following applies: (i) a response time shorter than 25 ns; (ii) a sample detection area larger than 149 mm2; (iii) a multiplication factor larger than 1,3 × 103. (c) a response time shorter than 5 ns for detecting electrons or ions; (d) a sample detection area larger than 314 mm2 for detecting electrons or ions; (e) a multiplication factor larger than 4,0 × 107. | Expires on the following dates: (a) 21 July 2021 for medical devices and monitoring and control instruments; (b) 21 July 2023 for in-vitro diagnostic medical devices; (c) 21 July 2024 for industrial monitoring and control instruments. |  |  | Active | 8 and 9 |  | Lead | 00021 | 2014-7-22 | Multiple | 2011/65/EU | 2014/70/EU | 2019-11-30 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01043-A-01 | 01043 | A | 01 | 39 | Others; Lead in micro-channel plates (MCPs) used in equipment where at least one of the following properties is present: (a) a compact size of the detector for electrons or ions, where the space for the detector is limited to a maximum of 3 mm/MCP (detector thickness + space for installation of the MCP), a maximum of 6 mm in total, and an alternative design yielding more space for the detector is scientifically and technically impracticable; (b) a two-dimensional spatial resolution for detecting electrons or ions, where at least one of the following applies: (i) a response time shorter than 25 ns; (ii) a sample detection area larger than 149 mm2; (iii) a multiplication factor larger than 1,3 × 103. (c) a response time shorter than 5 ns for detecting electrons or ions; (d) a sample detection area larger than 314 mm2 for detecting electrons or ions; (e) a multiplication factor larger than 4,0 × 107. | Expires on the following dates: (a) 21 July 2021 for medical devices and monitoring and control instruments; |  |  | Active | 8 and 9 other than in vitro and industrial |  | Lead | 00021 | 2014-7-22 | TBD | 2011/65/EU | 2014/70/EU | 2019-11-30 | 2021-08-23 | E3.0: Expiry Date extended (TBD); Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01043-A-02 | 01043 | A | 02 | 39 | Others; Lead in micro-channel plates (MCPs) used in equipment where at least one of the following properties is present: (a) a compact size of the detector for electrons or ions, where the space for the detector is limited to a maximum of 3 mm/MCP (detector thickness + space for installation of the MCP), a maximum of 6 mm in total, and an alternative design yielding more space for the detector is scientifically and technically impracticable; (b) a two-dimensional spatial resolution for detecting electrons or ions, where at least one of the following applies: (i) a response time shorter than 25 ns; (ii) a sample detection area larger than 149 mm2; (iii) a multiplication factor larger than 1,3 × 103. (c) a response time shorter than 5 ns for detecting electrons or ions; (d) a sample detection area larger than 314 mm2 for detecting electrons or ions; (e) a multiplication factor larger than 4,0 × 107. | Expires on the following dates: (b) 21 July 2023 for in-vitro diagnostic medical devices; |  |  | Active | 8 in vitro |  | Lead | 00021 | 2016-7-22 | 2023-07-21 | 2011/65/EU | 2014/70/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01043-A-03 | 01043 | A | 03 | 39 | Others; Lead in micro-channel plates (MCPs) used in equipment where at least one of the following properties is present: (a) a compact size of the detector for electrons or ions, where the space for the detector is limited to a maximum of 3 mm/MCP (detector thickness + space for installation of the MCP), a maximum of 6 mm in total, and an alternative design yielding more space for the detector is scientifically and technically impracticable; (b) a two-dimensional spatial resolution for detecting electrons or ions, where at least one of the following applies: (i) a response time shorter than 25 ns; (ii) a sample detection area larger than 149 mm2; (iii) a multiplication factor larger than 1,3 × 103. (c) a response time shorter than 5 ns for detecting electrons or ions; (d) a sample detection area larger than 314 mm2 for detecting electrons or ions; (e) a multiplication factor larger than 4,0 × 107. | Expires on the following dates: (c) 21 July 2024 for industrial monitoring and control instruments. |  |  | Active | 9 industrial |  | Lead | 00021 | 2017-7-22 | 2024-07-21 | 2011/65/EU | 2014/70/EU | 2019-11-30 | 2021-08-23 | E3.0: Request for renewal 2020-01-18 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01044-A-00 | 01044 | A | 00 | 40 | Others; Lead in dielectric ceramic in capacitors for a rated voltage of less than 125 V AC or 250 V DC for industrial monitoring and control instruments. | Expires on: 31 December 2020. May be used after that date in spare parts for industrial monitoring and control instruments placed on the market before 1 January 2021. |  |  | Expired | 9 industrial |  | Lead | 00021 | 2017-07-22 | 2020-12-31 | 2011/65/EU | 2014/69/EU | 2019-11-30 | 2021-08-23 | E3.0: Exemption expired 2020-12-31 E2.0: Exemption Status update 2021-01-19 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01045-A-00 | 01045 | A | 00 | 41 | Others; Lead as a thermal stabiliser in polyvinyl chloride (PVC) used as base material in amperometric, potentiometric and conductometric electrochemical sensors which are used in in-vitro diagnostic medical devices for the analysis of blood and other body fluids and body gases. | Expires on: 31 December 2018. |  |  | Expired | 8 in vitro |  | Lead | 00021 | 2016-07-22 | 2018-12-31 | 2011/65/EU | 2015/573/EU | 2019-11-30 | 2021-01-19 | E2.0: Exemption replaced and expired 2018-12-31 In Annex IV to Directive 2011/65/EU, entry 41 is replaced by the following: See Description (Exemption Description) of IEC-ID 01045-B-00 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01045-B-00 | 01045 | B | 00 | 41 | Others; Lead as a thermal stabiliser in polyvinyl chloride (PVC) used as base material in amperometric, potentiometric and conductometric electrochemical sensors which are used in in-vitro diagnostic medical devices for the analysis of blood and other body fluids and body gases. | Expires on 31 March 2022 |  |  | Active | 8 in vitro |  | Lead | 00021 | 2020-03-05 | 2022-03-31 | 2011/65/EU | 2020/366/EU | 2021-01-19 | 2021-01-19 | E2.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01046-A-00 | 01046 | A | 00 | 42 | Others; Mercury in electric rotating connectors used in intravascular ultrasound imaging systems capable of high operating frequency (> 50 MHz) modes of operation. | Expires on: 30 June 2026 |  |  | Active | 8 and 9 other than in vitro and industrial |  | Mercury | 00029 | 2015-04-30 | 2026-06-30 | 2011/65/EU | 2021/884/EU | 2021-01-19 | 2022-01-31 | E4.0: Request for renewal E3.0: Request for renewal 2017-10-06 E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01047-A-00 | 01047 | A | 00 | 43 | Others; Cadmium anodes in Hersch cells for oxygen sensors used in industrial monitoring and control instruments, where sensitivity below 10 ppm is required. | Expires on: 15 July 2023 |  |  | Active | 9 industrial |  | Cadmium | 00010 | 2016-07-15 | 2023-07-15 | 2011/65/EU | 2016/1029/EU | 2021-01-19 | 2021-01-19 | E1.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01048-A-00 | 01048 | A | 00 | 44 | Others; Cadmium in radiation tolerant video camera tubes designed for cameras with a centre resolution greater than 450 TV lines which are used in environments with ionising radiation exposure exceeding 100 Gy/hour and a total dose in excess of 100kGy. | Expires on 31 March 2027 |  |  | Active | 9 industrial |  | Cadmium | 00010 | 2020-03-05 | 2027-03-31 | 2011/65/EU | 2020/364/EU | 2021-01-19 | 2021-01-19 | E2.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01049-A-00 | 01049 | A | 00 | 45 | Bis(2-ethylhexyl) phthalate (DEHP) in ion-selective electrodes applied in point of care analysis of ionic substances present in human body fluids and/or in dialysate fluids | Expires on: 21 July 2028 |  |  | Active | 8 and 9 |  | Bis(2-ethylhexyl) phthalate (DEHP) | 00038 | 2021-07-21 | 2028-07-21 | 2011/65/EU | 2021/1980/EU | 2022-01-31 | 2022-01-31 | E4.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01050-A-00 | 01050 | A | 00 | 46 | Bis(2-ethylhexyl) phthalate (DEHP) in plastic components in MRI detector coils. | Expires on: 1 January 2024 |  |  | Active | 8 and 9 |  | Bis(2-ethylhexyl) phthalate (DEHP) | 00038 | 2021-07-21 | 2024-01-01 | 2011/65/EU | 2021/1979/EU | 2022-01-31 | 2022-01-31 | E4.0: Added exemption
IEC62474 | EU-RoHS-AnnexIV | E4.0 | 01051-A-00 | 01051 | A | 00 | 47 | Bis(2-ethylhexyl) phthalate (DEHP), butyl benzyl phthalate (BBP), dibutyl phthalate (DBP) and diisobutyl phthalate (DIBP) in spare parts recovered from and used for the repair or refurbishment of medical devices, including in vitro diagnostic medical devices, and their accessories, provided that the reuse takes place in auditable closed-loop business-to-business return systems and that each reuse of parts is notified to the customer. | Expires on: 21 July 2028 |  |  | Active | 8 and 8 in vitro |  | Bis(2-ethylhexyl) phthalate (DEHP), butyl benzyl phthalate (BBP), dibutyl phthalate (DBP) and diisobutyl phthalate (DIBP) | 00038; 00039; 00040; 00041 | 2021-07-21 | 2028-07-21 | 2011/65/EU | 2021/1978/EU | 2022-01-31 | 2022-01-31 | E4.0: Added exemption

## README
READ ME for Exemption Lists posted on the IEC 62474 Database
Important Information: With regard to legal relevance to the EU RoHS Directive:
The regulatory basis of this exemption list is Anenx III of the EU RoHS Directive 2011/65/EU and it's amendments. With regard to the EU RoHS Directive, this exemption list posted on the IEC 62474 Database is not legally binding and is for references purposes only to simply the use of exemptions by product manufacturers, suppliers, solution providers and other stakeholders. The IEC 62474 exemption lists are actively updated by the IEC 62474 database validation team (VT62474) but are provided as-is on a best efforts basis and have no legal representation for the EU RoHS Directive. Users should always refer to the exemptions in officially published Regulations, Directives, and Delegated Directives (as appropriate) for assessing compliance.
Important Information: What if another regulation references this exemption list in it's regulatory text
A regulation may reference this IEC 62474 International standard exemption list as a normative reference for exemptions allowed under the regulation. For example, a regulation may wish to reference IEC 62474 exemptions lists directly to enable on-going alignment with the exemptions in the EU RoHS Directive without the need to continuously amend the regulation. In such cases, this exemption list functions as any other IEC standard that is normatively referenced by a regulation and may be legally binding. Please refer to the official regulatory text for details.
Background Information on IEC 62474 Exemption Lists
The IEC 62474:2018 (published November 2018) includes a provision to post and maintain substance exemption lists on the IEC 62474 database. The initial release of the EU RoHS Annex III exemptions will include all of the exemption renewals that were published in 2018 and the additional ten Delegated Directives published in February 2019.
Identification of Exemption Lists
Each exemption list will be identified by the List authority = “IEC62474”, a list identity, and a list version. The list identity will remain fixed for a given regulation and the list version will be revised every time that an edit is made to the list. The list version is in the format Emmm.n where ‘mmm’ represents the major revision number and ‘n’ represents a minor editorial update.
Identification of Individual Entries in the Exemption List
Each exemption list entry will have a unique ID consisting of three parts: a base ID, a version ID, and a sub-exemption identity (explained below). This IEC-ID is intended to uniquely and permanently identify a specific exemption even when a regulation re-uses an exemption number after significantly modifying the technical aspects or scope of an exemption. The IEC-ID for each entry is of the format “nnnnn-vv-ss”, where nnnnn = base ID number, vv = version (alphabetic), and ss = sub-exemption ID (sub-identity).
Information about the column titled "Expiry Date"
A specific expiry date is provided when provided in the regulation. The following legend is applicable: ▪ yyyy-mm-dd: year, month and day in ISO date format ▪ 9999-12-31: this is a special-use expiry date indicating that the regulation does not specify an expiry and that the exemption period is considered indefinite ▪ TBD: "to be determined" - the expiry date is unknown -- this may occur when an exemption is under revision (e.g. a renewal request) and the applicable authority has not yet specified the new expiry date ▪ Multiple: the expiry date depends on the product or other factors associated with the exemption
Sub-exemptions: Exemption Details that Vary over Time or by Product Category
A unique aspect of the IEC 62474 exemption lists is that exemptions with technical requirements that change over time or have different effective/expiry dates based on product category will be subdivided into sub-exemptions that each have a unique technical requirement, effective date and/or expiry date. Each sub-exemption will share the base ID and version ID of its parent exemption but will have a unique sub-exemption identity (referred to as the sub-identity). For example, the revised exemption 6(a) that comes into effect on July 1, 2019 is excerpted in Table 6. The example lists three sub-exemptions with different expiry dates for most category 8/9 products, category 8 in vitro, and category 9 industrial. The IEC-IDs for the three sub-exemptions are: 00033-B-01, 00033-B-02, 00033-B-03 – the last two digits indicate the sub-identity. The original exemption as specified in the regulation always has the sub-identity ‘00’.